Supplementary data for

Quadrupole-Linear Ion Trap Tandem Mass Spectrometry System for Clinical Biomarker Analysis

Xiang Fang a,#,*, Jie Xie a,#, Shiying Chu a, You Jiang a, Yuting An a,b, Chang Li a,b, Xiaoyun Gong a, Rui Zhai a, Zejian Huang a, Chunling Qiu b, Xinhua Dai a,*

a Technology Innovation Center of Mass Spectrometry for State Market Regulation, Center for Advanced Measurement Science, National Institute of Metrology, Beijing 100029, China

b College of Instrument Science and Electrical Engineering, Jilin University, Changchun 130026, China

* Corresponding authors.

E-mail addresses: fangxiang@nim.ac.cn (X. Fang), daixh@nim.ac.cn (X. Dai).

# These authors contributed equally to this work.

Fig. S1. Mass spectrogram of ions which directly implanted into ion trap without quadrupole screening. (a) Creatine (CRE), (b) fragment ion of CRE, (c) D3-CRE, and (d) fragment ion of D3-CRE. m/z: mass-to-charge ratio; M: molecular weight..

Fig. S2. Mass spectrogram of ions which implanted into ion trap after screening by quadrupole. (a) CRE, (b) fragment ion of CRE, (c) D3-CRE, and (d) fragment ion of D3-CRE.

Fig. S3. Optimization of mass parameters. (a) Tube lens, (b) Q0, (c) PreQ, and (d) ROD. Q0: direct current bias voltage of ion guide Q0; PreQ: direct current bias voltage of pre-quadrupole; ROD: direct current bias voltage of rod.

Table S1

Quadrupole-linear ion trap (Q-LIT) mass parameters including quadrupole radio frequency (QRF) and ion trap radio frequency (ITRF) for GAA, D2-GAA, CRE, and D3-CRE.

| Compound | Ionization mode (ESI) | Precursor ions (m/z) | Product ions (m/z) | QRF (V) | ITRF (V) |
| --- | --- | --- | --- | --- | --- |
| GAA | M + H | 118.11 | 101.12 | 383.4 | 15 |
| D2-GAA | M + H | 120.09 | 103.12 | 390.1 | 15 |
| CRE | M + H | 132.10 | 90.11 | 429.6 | 15 |
| D3-CRE | M + H | 135.07 | 93.10 | 440.0 | 11 |

GAA: guanidinoacetic acid; CRE: creatine; ESI: electrospray ionization.

Table S2

Flow rate gradient condition.

| #### Time #### (min) | #### Flow rate #### (mL·min−1) | #### Sample volume #### (μL) | #### Mobile phase B #### (%) |
| --- | --- | --- | --- |
| 0 | 0.220 | 5.0 | 100.0 |
| 0.100 | 0.220 | 5.0 | 100.0 |
| 0.200 | 0.010 | 5.0 | 100.0 |
| 0.950 | 0.010 | 5.0 | 100.0 |
| 1.100 | 0.700 | 5.0 | 100.0 |
| 1.450 | 0.700 | 5.0 | 100.0 |
| 1.500 | 0.220 | 5.0 | 100.0 |

## Table S3

## Comparison FIA to LC method.

| Analytes | Matrix | Derivatization | Instrumental method | Separation column | Determination time (min) | Reference |
| --- | --- | --- | --- | --- | --- | --- |
| GAA | Plasma | No | HPLC–MS/MS | Hypersil gold C18 column (5 μm, 150 mm × 4.6 mm) | 8.0 | [1] |
| GAA CRE | Urine | No | LC–MS/MS | Xterra C18 column (3.5 μm, 2.1 mm × 30 mm) | 10.0 | [2] |
| GAA CRE Creatinine Ornithine arginine | Plasma Whole blood DBS | Yes 100 μL butanol–HCl (3 mol·L−1) incubated at 60 °C for 30 min | LC–MS/MS | Phenomenex Kinetex XB-C18 column (2.6 μm, 100 mm × 4.6 mm) | 10.0 | [3] |
| GAA CRE | Plasma | Yes Butanolic–HCl incubated at 65 °C for 15 min | LC–MS/MS | Supelcosil column (8.3 μm, 7.5 cm × 4.6 mm) | 5.0 | [4] |
| GAA CRE Creatinine | Urine Plasma | No | LC–MS/MS | C18 Synergi Fusion RP column (4 μm, 150 mm × 3.0 mm) | 2.5 | [5] |
| GAA CRE | Urine Plasma | Yes 3 mol·L−1 HCl in n-butanol for 15 min at 65 °C | LC–MS/MS | Symmetry C18 column (3.5 μm, 50 mm × 2.1 mm) | 3.0 | [6] |
| GAA CRE | Plasma | Yes Butanol/acetylchloride (90:10, v/v) incubated at 65 °C for 20 min | LC–MS/MS | Symmetry C18 column (3.5 μm, 50 mm × 2.1 mm) | 3.0 | [7] |
| GAA CRE Creatinine | Serum Urine | No | LC–APCI–MS | Inertsil ODS-2 packed column (5 μm, 250 mm × 4.6 mm) | 4.0 | [8] |
| GAA CRE | DBS | Yes 3 mol·L−1 hydrochloric acid in n-butanol solution at 60 °C for 30 min | FIA–ESI-MS/MS | None | 1.0 | [9] |
| GAA CRE | DBS | Yes Butanolic HCl for 20 min at 65 °C | FIA–ESI-MS/MS | None | - | [10] |
| GAA CRE Amino acids Acylcarnitines Succinylacetone | DBS | No | FIA–ESI-MS/MS | None | - | [11] |
| GAA CRE | DBS | No | FIA–ESI-Q-LIT | None | 1.5 | This work |

DBS: dried blood spot; FIA: flow injection analysis; LC: liquid chromatography; HPLC: high performance liquid chromatography; MS: mass spectrometry.

## Table S4

## QCs, linearity range, R2, LODs, LOQs, mean recoveries, and CV of CRE and GAA using Q-LIT V2 (National Institute of Metrology, China) and TSQ Altis (Thermo Fisher Scientific Inc., USA).

| Analytes | Instrument | QC-LL (μmol·L−1) | CV (%) (n = 10) | QC-HL (μmol·L−1) | CV (%) (n = 10) | Linearity range (μmol·L−1) | R2 | LOD (nmol·L−1) | LOQ (nmol·L−1) | Theoretical concentration (μmol·L−1) | Recovery (%) | CV (%) (n = 6) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| GAA | Q-LIT V2 | 5.06 | 6.38 | 9.21 | 6.88 | 1.20–9.84 | 0.9908 | 0.35 | 1.00 | 3.0 | 116.4 | 7.32 |
|  |  |  |  |  |  |  |  |  |  | 8.0 | 112.4 | 2.90 |
|  | TSQ Altis | 4.14 | 3.15 | 8.76 | 5. 42 | 1.20–9.56 | 0.9976 | 0.45 | 1.25 | 3.0 | 80.90 | 5.11 |
|  |  |  |  |  |  |  |  |  |  | 8.0 | 84.00 | 4.93 |
| CRE | Q-LIT V2 | 247.48 | 8.19 | 448.60 | 8.40 | 122.05–656.76 | 0.9902 | 0.07 | 0.20 | 150.0 | 114.6 | 3.19 |
|  |  |  |  |  |  |  |  |  |  | 300.0 | 119.9 | 7.73 |
|  | TSQ Altis | 198.68 | 5.36 | 426.15 | 4.78 | 125.51–658.27 | 0.9973 | 0.09 | 0.25 | 150.0 | 85.70 | 6.93 |
|  |  |  |  |  |  |  |  |  |  | 300.0 | 87.10 | 5.67 |

Range of QC-LL: 4.05–5.32 and 185.92–273.21 μmol·L−1 for GAA and CRE, respectively.

Range of QC-HL: 7.89–11.93 and 368.76–592.79 μmol·L−1 for GAA and CRE, respectively.

QC: quality control; R2: linear correlation coefficient; LOD: limit of detection; LOQ: limit of quantification; CV: coefficient of variation; QC-LL: low-level quality control; QC-HL: high-level quality control.

n is parallel experiments of the same sample.

## Table S5

## TSQ Altis mass parameters including RF lens and collision energy.

| Compounds | Ionization mode (ESI) | Precursor ions (m/z) | Product ions (m/z) | RF lens (V) | Collision energy (V) |
| --- | --- | --- | --- | --- | --- |
| GAA | M + H | 118.1 | 101.1 | 42 | 10.99 |
| D2-GAA | M + H | 120.0 | 103.1 | 42 | 10.23 |
| CRE | M + H | 132.0 | 90.2 | 52 | 14.28 |
| D3-CRE | M + H | 135.0 | 93.2 | 51 | 13.18 |

RF: radio frequency

Table S6

Determination of 77 actual clinical samples using Q-LIT V2 (NIM) and TSQ Altis (n = 5).

| No. | GAA |  |  |  |  |  | CRE |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Q-LIT V2 |  |  | TSQ Altis |  |  | Q-LIT V2 |  |  | TSQ Altis |  |
|  | Mean | CV |  | Mean | CV |  | Mean | CV |  | Mean | CV |
|  | (μmol·L−1) | (%) |  | (μmol·L−1) | (%) |  | (μmol·L−1) | (%) |  | (μmol·L−1) | (%) |
| 1 | 0.78 | 19.23 |  | 0.87 | 17.32 |  | 295.36 | 12.1 |  | 288.85 | 10.46 |
| 2 | 0.88 | 12.01 |  | 0.82 | 9.49 |  | 238.85 | 8.20 |  | 230.65 | 9.57 |
| 3 | 0.95 | 18.75 |  | 0.99 | 7.05 |  | 226.12 | 6.58 |  | 221.85 | 5.77 |
| 4 | 0.92 | 19.97 |  | 0.90 | 13.34 |  | 296.33 | 13.63 |  | 284.17 | 8.54 |
| 5 | 1.30 | 14.78 |  | 1.32 | 6.06 |  | 292.01 | 6.14 |  | 291.31 | 7.96 |
| 6 | 1.07 | 7.87 |  | 1.14 | 15.36 |  | 269.48 | 12.49 |  | 269.43 | 7.55 |
| 7 | 0.98 | 16.68 |  | 0.90 | 17.77 |  | 322.66 | 5.65 |  | 305.98 | 11.00 |
| 8 | 1.12 | 16.39 |  | 1.20 | 5.83 |  | 331.98 | 8.27 |  | 313.13 | 8.2 |
| 9 | 1.54 | 13.35 |  | 1.67 | 11.39 |  | 270.95 | 3.33 |  | 258.52 | 12.71 |
| 10 | 2.05 | 16.60 |  | 2.15 | 9.76 |  | 319.43 | 9.92 |  | 300.25 | 16.68 |
| 11 | 1.12 | 11.64 |  | 1.20 | 12.51 |  | 310.12 | 10.61 |  | 294.80 | 7.08 |
| 12 | 1.33 | 13.64 |  | 1.43 | 15.44 |  | 368.56 | 16.36 |  | 342.28 | 11.46 |
| 13 | 1.36 | 14.60 |  | 1.35 | 8.13 |  | 333.59 | 12.91 |  | 327.51 | 6.24 |
| 14 | 1.22 | 8.74 |  | 1.33 | 13.50 |  | 287.27 | 8.41 |  | 268.52 | 13.27 |
| 15 | 1.06 | 14.32 |  | 1.16 | 11.22 |  | 294.75 | 12.2 |  | 267.79 | 15.02 |
| 16 | 1.25 | 15.36 |  | 1.13 | 13.28 |  | 402.97 | 6.94 |  | 370.69 | 9.74 |
| 17 | 1.17 | 19.84 |  | 1.31 | 17.51 |  | 271.10 | 7.20 |  | 253.62 | 7.27 |
| 18 | 1.55 | 19.93 |  | 1.63 | 8.58 |  | 310.26 | 10.18 |  | 289.20 | 13.24 |
| 19 | 1.50 | 17.21 |  | 1.54 | 9.61 |  | 364.41 | 3.36 |  | 332.78 | 14.46 |
| 20 | 1.10 | 13.12 |  | 1.05 | 14.29 |  | 357.31 | 11.15 |  | 334.98 | 8.49 |
| 21 | 1.29 | 17.82 |  | 1.34 | 17.15 |  | 362.36 | 12.42 |  | 357.96 | 10.85 |
| 22 | 1.09 | 11.94 |  | 1.16 | 19.00 |  | 284.37 | 10.11 |  | 277.09 | 7.01 |
| 23 | 1.50 | 19.12 |  | 1.47 | 8.16 |  | 341.47 | 7.44 |  | 350.04 | 4.57 |
| 24 | 1.55 | 18.90 |  | 1.69 | 13.61 |  | 341.60 | 11.56 |  | 350.39 | 12.85 |
| 25 | 1.41 | 15.90 |  | 1.55 | 18.09 |  | 359.68 | 5.40 |  | 362.21 | 2.27 |
| 26 | 1.68 | 17.11 |  | 1.72 | 18.64 |  | 327.70 | 4.33 |  | 341.33 | 4.34 |
| 27 | 1.70 | 18.53 |  | 1.87 | 17.67 |  | 318.15 | 5.76 |  | 322.13 | 16.32 |
| 28 | 1.51 | 18.75 |  | 1.63 | 13.51 |  | 290.76 | 9.38 |  | 286.73 | 5.09 |
| 29 | 1.63 | 16.01 |  | 1.78 | 11.81 |  | 398.56 | 7.66 |  | 388.72 | 2.72 |
| 30 | 2.06 | 15.93 |  | 2.03 | 6.06 |  | 580.28 | 12.83 |  | 575.04 | 15.5 |
| 31 | 1.82 | 16.68 |  | 1.74 | 4.04 |  | 380.79 | 8.99 |  | 368.45 | 5.19 |
| 32 | 2.70 | 3.31 |  | 2.77 | 9.09 |  | 267.42 | 5.04 |  | 258.34 | 4.98 |
| 33 | 2.02 | 5.12 |  | 2.04 | 7.17 |  | 276.65 | 7.71 |  | 275.16 | 9.11 |
| 34 | 1.55 | 11.7 |  | 1.51 | 2.45 |  | 211.72 | 14.97 |  | 205.17 | 10.01 |
| 35 | 1.86 | 15.07 |  | 1.84 | 5.60 |  | 259.09 | 8.68 |  | 265.57 | 9.21 |
| 36 | 1.57 | 18.57 |  | 1.67 | 8.32 |  | 268.69 | 7.66 |  | 275.05 | 9.46 |
| 37 | 2.29 | 16.63 |  | 2.37 | 8.10 |  | 219.01 | 12.22 |  | 209.92 | 12.76 |
| 38 | 1.79 | 12.87 |  | 1.76 | 4.07 |  | 283.34 | 13.80 |  | 282.03 | 6.37 |
| 39 | 1.43 | 7.33 |  | 1.46 | 8.21 |  | 382.85 | 8.34 |  | 400.19 | 6.26 |
| 40 | 2.23 | 11.88 |  | 2.41 | 10.77 |  | 515.99 | 10.91 |  | 490.71 | 5.87 |
| 41 | 2.31 | 16.91 |  | 2.31 | 4.17 |  | 580.33 | 9.42 |  | 561.70 | 7.73 |
| 42 | 2.26 | 11.65 |  | 2.52 | 16.06 |  | 349.74 | 11.33 |  | 341.55 | 4.09 |
| 43 | 1.97 | 9.91 |  | 2.16 | 12.04 |  | 297.18 | 8.24 |  | 290.67 | 8.77 |
| 44 | 2.07 | 17.49 |  | 2.27 | 9.45 |  | 371.34 | 5.28 |  | 377.76 | 4.87 |
| 45 | 2.53 | 9.36 |  | 2.38 | 12.23 |  | 299.46 | 10.03 |  | 298.09 | 10.16 |
| 46 | 2.06 | 17.33 |  | 2.05 | 3.44 |  | 259.34 | 7.21 |  | 268.56 | 6.96 |
| 47 | 1.63 | 14.74 |  | 1.55 | 11.22 |  | 431.70 | 3.36 |  | 438.25 | 2.34 |
| 48 | 2.11 | 18.37 |  | 2.19 | 8.68 |  | 370.44 | 7.22 |  | 372.88 | 4.42 |
| 49 | 2.16 | 17.28 |  | 2.11 | 3.26 |  | 285.92 | 9.39 |  | 299.70 | 8.05 |
| 50 | 2.49 | 18.83 |  | 2.61 | 9.64 |  | 357.95 | 7.86 |  | 365.43 | 7.00 |
| 51 | 2.03 | 13.49 |  | 2.06 | 14.43 |  | 238.27 | 6.32 |  | 234.04 | 2.01 |
| 52 | 1.80 | 18.34 |  | 1.66 | 9.77 |  | 253.90 | 11.97 |  | 240.86 | 6.43 |
| 53 | 2.02 | 13.99 |  | 1.93 | 8.94 |  | 340.02 | 13.28 |  | 325.40 | 6.83 |
| 54 | 1.94 | 14.98 |  | 2.03 | 4.27 |  | 290.29 | 9.89 |  | 287.09 | 1.68 |
| 55 | 1.79 | 12.23 |  | 1.81 | 7.21 |  | 479.85 | 6.29 |  | 481.94 | 1.91 |
| 56 | 1.89 | 16.30 |  | 2.17 | 17.34 |  | 358.63 | 8.97 |  | 355.69 | 5.13 |
| 57 | 1.97 | 13.65 |  | 1.85 | 7.70 |  | 334.94 | 8.47 |  | 330.81 | 8.26 |
| 58 | 1.37 | 14.49 |  | 1.41 | 12.39 |  | 367.19 | 8.42 |  | 351.55 | 9.44 |
| 59 | 1.98 | 17.65 |  | 2.01 | 5.87 |  | 257.93 | 8.80 |  | 251.20 | 6.54 |
| 60 | 1.32 | 9.07 |  | 1.37 | 12.05 |  | 316.50 | 6.14 |  | 309.92 | 7.13 |
| 61 | 1.71 | 19.76 |  | 1.72 | 6.91 |  | 401.76 | 10.15 |  | 391.92 | 7.98 |
| 62 | 1.53 | 17.03 |  | 1.48 | 14.84 |  | 386.58 | 8.10 |  | 375.27 | 10.58 |
| 63 | 1.65 | 19.38 |  | 1.75 | 9.71 |  | 272.14 | 11.29 |  | 253.24 | 9.45 |
| 64 | 1.29 | 13.17 |  | 1.30 | 7.87 |  | 293.70 | 10.6 |  | 286.58 | 7.31 |
| 65 | 1.51 | 12.51 |  | 1.59 | 8.79 |  | 509.35 | 9.23 |  | 484.71 | 6.61 |
| 66 | 1.70 | 13.45 |  | 1.76 | 8.42 |  | 198.68 | 9.92 |  | 182.44 | 9.23 |
| 67 | 1.85 | 15.52 |  | 1.94 | 7.10 |  | 319.82 | 7.08 |  | 295.43 | 9.94 |
| 68 | 1.24 | 19.04 |  | 1.27 | 3.61 |  | 214.48 | 12.5 |  | 210.01 | 5.75 |
| 69 | 1.12 | 13.41 |  | 1.17 | 11.09 |  | 248.55 | 13.27 |  | 235.77 | 6.71 |
| 70 | 1.39 | 18.72 |  | 1.33 | 7.19 |  | 294.27 | 11.53 |  | 278.13 | 12.62 |
| 71 | 1.54 | 11.98 |  | 1.48 | 9.78 |  | 219.82 | 8.84 |  | 220.81 | 4.21 |
| 72 | 1.28 | 17.21 |  | 1.17 | 12.47 |  | 233.86 | 13.26 |  | 239.74 | 12.09 |
| 73 | 1.34 | 9.74 |  | 1.25 | 9.56 |  | 282.18 | 6.62 |  | 287.37 | 6.52 |
| 74 | 1.31 | 18.6 |  | 1.20 | 10.93 |  | 337.71 | 12.24 |  | 334.99 | 13.88 |
| 75 | 1.47 | 9.54 |  | 1.42 | 6.64 |  | 194.63 | 6.86 |  | 190.45 | 6.19 |
| 76 | 1.06 | 17.95 |  | 1.16 | 12.25 |  | 217.80 | 11.54 |  | 211.13 | 10.41 |
| 77 | 0.98 | 10.82 |  | 1.01 | 6.90 |  | 255.97 | 10.27 |  | 239.87 | 9.60 |

References

[1] Sun N, Wu Y, Zhao L, He H, Mei D, Zhang S, et al. A rapid and sensitive HPLC–MS/MS method for determination of endogenous creatine biosynthesis precursors in plasma of children with viral myocarditis. J Chromatogr B 2019;1118–1119:148–56.

[2] Benoit R, Samir M, Boutin J, Samuel A, Brigitte C, Dominique D, et al. LC–MS/MS measurements of urinary guanidinoacetic acid and creatine: method optimization by deleting derivatization step. Clin Chim Acta 2019;493:148–55.

[3] Tran C, Yazdanpanah M, Kyriakopoulou L, Levandovskiy V, Zahid H, Naufer A, et al. Stable isotope dilution microquantification of creatine metabolites in plasma, whole blood and dried blood spots for pharmacological studies in mouse models of creatine deficiency. Clin Chim Acta 2014;436:160–8.

[4] Boenzi S, Rizzo C, Di Ciommo VM, Martinellia D, Goffredo BM, Marca G, et al. Simultaneous determination of creatine and guanidinoacetate in plasma by liquid chromatography–tandem mass spectrometry (LC–MS/MS). J Pharmaceut Biomed 2011;56(4):792–8.

[5] Carling RS, Hogg SL, Wood TC, Calvin J. Simultaneous determination of guanidinoacetate, creatine and creatinine in urine and plasma by un-derivatized liquid chromatography–tandem mass spectrometry. Ann Clin Biochem 2008;45(6):575–84.

[6] Cognat S, Cheillan D, Piraud M, Roos B, Jakobs C, Vianey-Saban C. Determination of guanidinoacetate and creatine in urine and plasma by liquid chromatography–tandem mass spectrometry. Clin Chem 2004;50(8):1459–61.

[7] Bodamer OA, Bloesch SM, Gregg AR, Stockler-Ipsiroglu S, O’Brien WE. Analysis of guanidinoacetate and creatine by isotope dilution electrospray tandem mass spectrometry. Clin Chim Acta 2001;308(1–2):173–8.

[8] Yasuda M, Sugahara K, Zhang J, Ageta T, Nakayama T, Shuin T, et al. Simultaneous determination of creatinine, creatine, and guanidinoacetic acid in human serum and urine using liquid chromatography–atmospheric pressure chemical ionization mass spectrometry. Anal Biochem 1997;253(2):231–5.

[9] Carducci C, Santagata S, Leuzzi V, Carducci C, Artiola C, Giovanniello T, et al. Quantitative determination of guanidinoacetate and creatine in dried blood spot by flow injection analysis–electrospray tandem mass spectrometry. Clin Chim Acta 2006;364(1–2):180–7.

[10] Pasquali M, Schwarz E, Jensen M, Yuzyuk T, DeBiase I, Randall H, et al. Feasibility of newborn screening for guanidinoacetate methyltransferase (GAMT) deficiency. J Inherit Metab Dis 2014;37(2):231–6.

[11] Asef C, Khaksarfard KM, De Jesús VR. A non-derivatized assay for the simultaneous detection of amino acids, acylcarnitines, succinylacetone, creatine, and guanidinoacetic acid in dried blood spots via tandem mass spectrometry. Int J Neonatal Screen 2016;2(4):13.